## Slide 1: ODS Next Generation
Technical Committee Meeting
June 4, 2019

## Slide 2: ODS Next Generation
Background
Value and Purpose of the System
Next Gen Workgroup Evaluation
Summary Of Findings
Next Steps

## Slide 3: ODS Background
1978 - Monitoring system developed in response to major carbon tetrachloride spill on Kanawha River
Two key functions:
  Spill detection
  Spill tracking
17 stations
Three system types
  GC/MS
  GC/FID
  Process GC
Quantify 30 VOC analytes
Detect thousands of VOCs

## Slide 4: Volatile Organics Monitoring Network
Notes: Our network is currently comprised of 17 active monitoring stations: by Instrumentation we have 9 GCMS, 3 GC(FID) 5 GC(AID) (on-line process GC’s). Our newest site, on the Elk River, near Charleston became active in Fall 2017.

## Slide 5: Value of System
Serves as a sentinel to alert water utilities
Examples
1994 – EDB from Shell Chemical fire
2009 – Unreported methylene chloride release
2014 – MCHM release in Charleston, WV
2017 – Parkersburg warehouse fire
Quick turn-around screening provides water utilities with key information to make informed management decisions to protect consumers

## Slide 6
Thursday, Jan 9th, 1200
Sunday, Jan 12th, 2139
Monday, Jan 13th , 0300
Monday, Jan 13th, 1400
Wednesday, Jan 16th, 0700
Tuesday, Jan 15th, 0600
Friday, Jan 17th, 0530
Sunday, Jan 19th, 2330
Thursday, Jan 23rd, 1000

## Slide 7
Notes: Based on Jar testing testing, we were comfortable closing the intakes when the raw water concentration reached 2 ug/l, No detectable EDB made it to the GAC. All of it was removed with the PAC

## Slide 8: ODS Next Generation
Previous renovation initiated in 2009
Work Group Participants
  American Water Company
  Greater Cincinnati Water Works
  Louisville Water Company
  Northern Kentucky Water District
  US EPA
  ORSANCO
Work Group Charge –
  Evaluate potential contaminants of concern
  Evaluate available technologies
  Identify at least 3 monitoring designs
  Recommend preferred option

## Slide 9: Process
Online Surveys:
  Polled Ohio River water utilities to determine system needs and areas for improvement
Researched:
  Contaminants of concern
  New instrument/sensor technologies
System Configuration Options –
  Developed list of system design options
Scoring Matrix –
  Developed matrix to score sites for GC/MS placement

## Slide 10: Online Survey Results
Great participation
  Responses from 27 utilities
Key Findings
  Significant variability among utilities
    Size, expertise, monitoring resources
  Most desired system enhancements
    Host water quality data web portal
    Broaden analytical capabilities

## Slide 11: Key Contaminants
VOCs (expand list?)
SVOCs, HAB indicators, Inorganic anion/cations, nutrients
SVOCs most beneficial; however, costly and labor intensive
Utilities routinely collect basic water quality parameters

## Slide 12: Instrument Evaluation
Work Group discussed various instrument/monitoring options
Alternate sensor types lack necessary detection limits.
HAP unit (portable GC) offers portability; however, is more expensive than current GC/MS system
Basic water quality parameter monitoring system insufficient to detect spills

## Slide 13
ODS System Configuration Scenarios |  |  |  |  |  |  | 
 |  |  |  | Annualized Cost |  |  | 
# | Description | Instrumentation | Total Capital | Maintenance Labor Hours | Capital Cost* | Maintenance Cost** | Total Annual Cost
1 | Current system + SVOC at 4 sites | 9 GC/MS 3 GC/FID 4 CMS 4 GC/FID (SVOC) 1 Backup CMS | $2,441,000 | 550 | $276,100 | $271,500 | $547,600
2 | Current system | 9 GC/MS 3 GC/FID 4 CMS 1 Backup CMS | $2,121,000 | 438 | $212,100 | $207,500 | $419,600
3 | Replace 3 GC/MS with GC/FID | 6 GC/MS 6 GC/FID 4 CMS 1 Backup CMS | $2,013,000 | 390 | $201,300 | $187,000 | $388,300
4 | Replace 5 GC/MS with 3 GC/FID & 2 CMS | 4 GC/MS 6 GC/FID 6 CMS 1 Backup CMS | $1,803,000 | 394 | $180,300 | $169,400 | $349,700
* Annualized capital cost assumes 10-year replacement schedule. |  |  |  |  |  |  | 
** Annualized maintenance cost excludes staff labor and travel expenses. |  |  |  |  |  |  | 

## Slide 14: Scoring Matrix
Developed matrix to score ODS sites for possible GC/MS placement
Scoring factors
Facility ODS performance metrics
Potential for spill occurrences
Enhanced monitoring capacity during spills
Population served
Wild Card - Geographic distribution

## Slide 15: Matrix Results
Prioritization of GC/MS Sites
Cincinnati owns their GC/MS; therefore, was excluded from scoring
Rank | Site | Composite Score
1 | Huntington | 1.4
2 (tied) | Louisville Wheeling | 1.5
4 (tied) | Evansville Hays Mine | 1.8

## Slide 16: Summary of Recommendations
Maintain VOCs as primary focus for routine monitoring system.
Expand calibrated suite of VOC analytes for routine analysis.
Scoring matrix developed can be used to prioritize placement of GC/MS instruments.

## Slide 17: Summary of Recommendations
Trial SVOC analysis at 2-3 sites with existing equipment to evaluate cost and time implications
Evaluate possible web-portal for utilities to share routinely-collected water quality results
Continue ODS Next Gen Work Group to guide ongoing recommendations.

## Slide 18: Summary of Recommendations
The three GC instrument system types currently employed remain the preferred instrument choices.
Four configuration options presented provides decision tool to optimize system to available resources.
Network with 5 GC/MS units considered minimum required to meet routine and spill response monitoring needs.
Maintaining current system configuration deemed preferred option.

## Slide 19: Questions?