## Slide 1: ODS, Spills & Emergency Response Update
Lila X. Ziolkowski, Travis Luncan
213th Technical Committee Meeting
February 7-8, 2017

## Slide 2: ODS Renovation Status
Project ended 09/30/16
  Acquired computers for upgrade at all sites.
  Purchased consumable supplies and parts for inventory,
  YSI Sonde for basic WQ parameters
  ODS Website undergoing minor fixes
Reimbursement requests #22, 23 approved by USEPA
  Received $139,000
  Maintenance agreement on GC and GCMS components through 12/31/18.
Notes: 1 yr maintenance package won’t go into effect until website is fully operational.
Since grant is over, access to website will be for utility personnel
Since 2016, staff has addressed all purge and trap issues and performed R&M on them.

## Slide 3: ODS Current Operations
Sites that still need computers to be installed:
(Wheeling, St. Albans, Portsmouth and Paducah) 12 have been installed.
GC/MS issues:
  Louisville site down, estimated that site will be back up the 2nd week of February.
    MS detector issues
    next downstream site notified
  ORSANCO HQ Mass Spec had multiple board failures, site now operational.
    Vendor fixed issue under service agreement.
CMS5000 issues:
    Ashland concentrator tube replacement on site by ODS staff.
Notes: Issues with Parts availability (from Singapore), partial instrument rebuild. Evansville notified.

## Slide 4
GC/FID: No issues reported
Purge and Trap auto sampler:
  Sent the Portsmouth Water auto sampler back to the vendor. Repair cost was covered by the vendor.
  Hays Mine needle drive had to be completely refurbished by ODS staff.
  Purge and Trap Concentrator Issues
  St. Albans, obstruction, overflow of water into the purge and trap cascaded into GC. ODS staff completed a full refurbishment of the purge and trap.
  Hays Mine obstruction in purge and trap was corrected with replacement of sample frit.
ODS Current Operations

## Slide 5: ODS Site Performance Figures September 1st – December 31st 2016 (122 days of ODS operations)
Daily river water samples are collected and analyzed by each of the ODS sites. This counts toward the daily utilization value. The graph below shows the average number of daily river water samples collected and analyzed and average number of days those samples were not collected and analyzed
Daily river water samples were not collected an average of 16 days
Daily river water samples were collected and analyzed an average of 106 days
Utilization % was (106 daily river water samples) X 100% =
122 days
87.2%

## Slide 6
ODS Site Performance Figures September 1st – December 31st 2016 (122 days of ODS operations)
A number of ODS sites collect and analyze multiple river water samples per day (4-6 samples). In the last four months of 2016, 6,469 samples were screened by our ODS system.
This represents an average of between 3 and 4 river water samples analyzed by each ODS site per day.
By providing additional training and staying
on top of instrumentation issues, staff ensures
good performance of the ODS in the future.

## Slide 7: Emergency Response
2017 Emergency Response Directory finalized
  Staff received Agency Use Copy
  Electronic public use copy available on ORSANCO website
  Printed copies available for distribution

## Slide 8: 2016 Spills Summary January 1st – December 31st
590 NRC (National Response Center) reports were received by ORSANCO spills staff in2016.
Of those 227 showed a potential impact to the Ohio River.
The top 12 materials listed on the NRC reports potentially impacting the Ohio River were
Material Name | # of Spills by Material | Material Name | # of Spills by Material
Unknown Oil | 92 | Motor Oil | 5
Hydraulic Oil | 34 | Coal Slurry | 3
Diesel Fuel | 26 | Fuel Oil | 3
Crude Oil | 8 | Paint | 3
Lube Oil | 8 | Bilge Slops | 3
Gasoline | 8 | Mineral Oil | 3